## Project Checklist Templates
Id | Type | Name | Values | Prompting Name
db55fcb1-ea39-4fe3-aa17-c5558f00f3e7 | checklist | Entire Village - Fire to Nurse Call interface test |  | Entire Village - Fire to Nurse Call interface test
a07f9a46-28d1-4ec0-b608-5fc1c4865971 | section | Details
ab373724-3bfd-4106-adf9-6d27218220be | checkpoint | Programme all Annunciators in the Village with text scrolling (Attach Photos)
e046a47f-f3e3-4c9b-b481-cf992d16eccb | checklist | Hospital - Display - Hospital 2 - Ollivier |  | Hospital - Display - Hospital 2 - Ollivier
4cc7d973-a5c0-4197-9fa2-3c203f9335a6 | section | Details
dd72b721-f02a-4c18-be20-33673a1388a0 | checkpoint | Check Displays in the Hallway near Room 228, one looks like it has no power (Attach Photos)
034fbb87-34cc-4482-87f8-ae1ea6764471 | checklist | Hospital - Hospital 2 - Room 203 |  | Hospital - Hospital 2 - Room 203
145b3ff7-94c7-474c-bc74-96682821a6b6 | section | Details
7ac0c2ea-7c7d-460d-95ca-128438c7ff56 | checkpoint | Check Staff Assist alarm, text description is wrong on the displays - shows "H208" (Attach Photos)
dc24f357-775b-4165-b10c-3752c7c8b20a | checkpoint | Check Floor Pressure mat, as it failed (Attach Photos)
5578439f-e21e-4120-aefc-20cec72d801e | checklist | Hospital - Hospital 2 - Room 210 |  | Hospital - Hospital 2 - Room 210
00eb76e1-7ce2-4da2-83f9-9aee12488318 | section | Details
21c58694-1aed-4777-b088-dffb6c758b99 | checkpoint | Check Staff Assist Alarm, as it failed (Attach Photos)
c76e93b5-311a-467b-b683-e5c8bd233ba1 | checkpoint | Check Floor Pressure mat, as it failed (Attach Photos)
7adee323-605b-4270-b4ff-3f95df5edf5f | checklist | Hospital - Hospital 2 - Room 212/219 |  | Hospital - Hospital 2 - Room 212/219
7b358bed-7363-41fe-bfc2-f2b7093aaa7a | section | Details
3bf4b3dd-548b-4478-a256-c8414c172ea8 | checkpoint | Replace Bedhead Call Point - requires very hard pressing to activate (Attach Photos)
8330eb53-a8a3-4d27-ae83-6533bb35e05c | checklist | Hospital - Hospital 2 - Room 223/242 |  | Hospital - Hospital 2 - Room 223/242
addac9d7-1ab3-4bef-a7c3-f0bcfe4cc161 | section | Details
31124ace-c90a-4b4e-8f89-0928ea5a93f9 | checkpoint | Check wired pendant lead, as it failed (Attach Photos)
4e308a79-9bd7-4858-93f2-c15af439df03 | checklist | Hospital - Hospital 2 - Room 225 |  | Hospital - Hospital 2 - Room 225
60db981c-05a9-410d-93fa-93e726d2fc0f | section | Details
238efbc5-3ed6-429d-bd5d-2705bc18a94f | checkpoint | Change Ensuite Toilet from Emergency Call Point to a Resident Call Point (Attach Photos)
92988095-144f-4ff1-aac0-a6b18f3fc857 | checklist | Hospital - Hospital 2 - Room 226 |  | Hospital - Hospital 2 - Room 226
7b1dca94-2890-4ec3-84c4-55d39551b51a | section | Details
e8311fe5-66c2-487a-bb57-7bf812d7ff3b | checkpoint | Change Floor Pressure Mat connection cable, as it is damaged (Attach Photos)
09f4af54-696a-42c5-8f9a-4f9409db3547 | checklist | Hospital - Hospital 2 - Room 235 |  | Hospital - Hospital 2 - Room 235
5312e3cc-3889-458b-8eca-b3043bbe3586 | section | Details
72e1e4c7-9cd0-4f45-ae10-6cabf08bc6e6 | checkpoint | Check ODL - illuminates Red for Staff Assist and/or Resident Call Alarm - should be Flashing Yellow for Staff Assist and Solid Yellow for Resident Call (Attach Photos)
1fc45381-e1e2-4fb1-9dd4-8ba93647d2bd | checkpoint | Check Red Bedhead Call Point, this is activating a Resident Call Alarm and should be activating an Emergency Call Alarm (Attach Photos)
d324e837-8b51-496b-a887-927782f73c78 | checklist | Hospital - Hospital 2 - Room 236 |  | Hospital - Hospital 2 - Room 236
a87acda3-02d3-4982-a6c2-326ed3d3303c | section | Details
3feaa2f9-0aad-404b-a174-f8f4b6a5e879 | checkpoint | Install surrounding cover on the ODL (Attach Photos)
d63b0154-d033-46b8-9929-aa85d82017e6 | checklist | Hospital - Hospital 2 - Room 237 |  | Hospital - Hospital 2 - Room 237
4f4cc29c-f288-4539-9055-600cb3690f04 | section | Details
1877c05c-cbbf-4573-95cb-d50a3f9dbde0 | checkpoint | Change Floor Pressure Mat connection cable, as it is damaged (Attach Photos)
9a61ef5e-930d-4159-a15f-efaa78f8e5a9 | checkpoint | Change Splitter cable, as it is damaged (Attach Photos)
99667022-051e-46be-80c2-0f5eb1273352 | checklist | Hospital - Hospital 2/3 - Room 207/208/220/451 |  | Hospital - Hospital 2/3 - Room 207/208/220/451
4906f2d6-3d69-44a0-920e-919c97f51474 | section | Details
de19d4ec-776b-46b5-8cc9-0b4f3cfa591b | checkpoint | Install surrounding cover on the Bedhead Call Point (Attach Photos)
ad71a28e-b390-4308-bc6e-a68abdc0681f | checklist | Hospital - Hospital 2/3/Independent Apartments - Tenzing - Room 221/238/435/442/447/452/455 |  | Hospital - Hospital 2/3/Independent Apartments - Tenzing - Room 221/238/435/442/447/452/455
a6b92828-97cb-416b-b0aa-0b716e1a6c87 | section | Details
6a13fe3a-3f5f-4544-8ec0-01005d264524 | checkpoint | Tidy wireless pendant cables (Attach Photos)
f41f7849-6f91-40e2-91fc-00d13fd9ba8b | checkpoint | Mount receiver onto the wall (Attach Photos)
3fe4a6cc-dc4c-4645-aff2-1cfbeeacb83b | checklist | Hospital - Hospital 3 - Room 434/458 |  | Hospital - Hospital 3 - Room 434/458
d81615f9-4dd2-4803-9a20-7218fafc6e39 | section | Details
920b0591-e09d-4fe0-928f-ac74403ebed0 | checkpoint | Install surrounding cover on the Ensuite Toilet Call Point (Attach Photos)
8aba4e03-1de0-4bb5-8e73-e4909c239cb0 | checklist | Hospital - Hospital 3 - Room 454 |  | Hospital - Hospital 3 - Room 454
34a9d0fa-6bae-4d92-9a46-4f59c1ef62b8 | section | Details
b56f7db2-ed68-4600-b921-4207493cf702 | checkpoint | Check Staff Assist alarm, text description is wrong on the displays - shows "H3 Rm455" (Attach Photos)
b5d05078-02ab-4820-98fb-5e23b3b48278 | checkpoint | Tidy wireless pendant cables (Attach Photos)
11b92a07-c44b-478d-94c1-a471ca19fae3 | checkpoint | Mount receiver onto the wall (Attach Photos)
27fcb5d7-f3d4-43ed-bb37-7551bbd596b3 | checklist | Hospital - Hospital 3 - Room 457 |  | Hospital - Hospital 3 - Room 457
d32f6297-429c-4d91-bb98-a88ccc05b473 | section | Details
9fd76842-cfdf-4ef3-9112-e33df34c4417 | checkpoint | Replace insert in the Ensuite Toilet Call Point, requires Green Insert (Attach Photos)
89126b7d-b3a8-43bd-8017-5e76001e295f | checklist | Hospital - Hospital 3 - Room 459 |  | Hospital - Hospital 3 - Room 459
d0ff635c-d08b-4342-988d-af9a07f9f173 | section | Details
022bc785-fc00-4dea-b914-ce2663d9ba81 | checkpoint | Install surrounding cover on the Bedhead Call Point (Attach Photos)
32889029-e8ea-4661-885d-6d1ece92132c | checkpoint | Tidy wireless pendant cables (Attach Photos)
f31a02ae-1c70-4a09-96bc-261364046810 | checkpoint | Mount receiver onto the wall (Attach Photos)
e4436cda-6cbf-4060-9b85-de58448b7ff7 | checklist | Hospital - Hospital 3 - Toilet |  | Hospital - Hospital 3 - Toilet
3f50c39c-9bc3-40a8-9f4a-ad8ac24b6b8e | section | Details
76000742-deda-4810-b505-2c06936034f2 | checkpoint | Install surrounding cover on the Toilet Call Point (Attach Photos)
1b8ea8c8-e8d5-4f01-9a97-9fcfc7f509c9 | checklist | Independent Apartment - Mallory/Shackleton - 121/433/146 |  | Independent Apartment - Mallory/Shackleton - 121/433/146
6cd2ff70-9b59-4fd7-9890-bcbd7130b864 | section | Details
5a130ea9-24b0-4ba3-8631-6987c65410c6 | checkpoint | Replace Shower Pull Cord, as it requires to be manually pushed back up to rest position to cancel the call (Attach Photos)
d261339b-2940-48b5-b19b-c7bf435c959d | checklist | Independent Apartments - Cook - 175 |  | Independent Apartments - Cook - 175
700ee815-c914-4052-8052-ac31ba3c5cf3 | section | Details
a4a2d0f0-3ae3-40fc-9d2d-38d73bdae9bc | checkpoint | Remove the Miracle RF Wireless Pendant Receiver that is installed, as the resident does not have a wireless pendant (Attach Photos)
a8f28478-9134-491e-a37b-0d687cf75aa9 | checklist | Independent Apartments - Cook - 176 |  | Independent Apartments - Cook - 176
a4a9c57d-3988-4077-8beb-7c153d6b079b | section | Details
049abddf-d232-41c7-a047-42eaf018df35 | checkpoint | Replace Call Socket on the Master Bedroom Bedhead Call Point, as It is broken (Attach Photos)
5669bb6d-e56a-4408-aa16-4d5708b98064 | checklist | Independent Apartments - Cook - 279 |  | Independent Apartments - Cook - 279
b61ea111-8974-4440-8193-73d019b97a7b | section | Details
05cb8606-d77c-4aa7-8e1d-22c9942375cc | checkpoint | Provide longer corded pendant, as resident advises current 1.5m pendant only just reaches across (Attach Photos)
6206ba4b-41fa-44a2-9c96-0f7af02de97a | checklist | Independent Apartments - Cook - Activities Female WC |  | Independent Apartments - Cook - Activities Female WC
5129884c-37f4-46e5-8a8b-df564bdd824e | section | Details
2581a247-98c2-4536-97e6-83cf385c77df | checkpoint | Check both entrances ODL's, as one failed (Attach Photos)
231d3257-1daf-49b0-baf5-aede07df0e5f | checklist | Independent Apartments - Everest - 163 |  | Independent Apartments - Everest - 163
ab0f61f0-675c-4cd8-b2da-573a6d9f26f5 | section | Details
5da6b14c-61f5-4363-88b5-befac1886eb1 | checkpoint | Check Master Bedroom Bedhead Pendant removal alarm, as it failed (Attach Photos)
a9ee38a2-4e09-4ffd-9af3-414dfad216cf | checklist | Independent Apartments - Everest - 263 |  | Independent Apartments - Everest - 263
ee18f21d-b079-420d-b5e8-332506470c20 | section | Details
3b9e1f5a-c2ba-412d-8a1d-db3d72a294eb | checkpoint | Check Master Bedroom Bedhead Pendant removal alarm, as it failed (Attach Photos)
f0dbd966-5695-4e64-b230-96b0c4852d7a | checkpoint | Check Guest Showers Pull Cord Call Point, as it is not securely mounted (Attach Photos)
34322ebc-f453-4dd4-8a9c-1ee5f7c1a3dc | checklist | Independent Apartments - Mallory - 122/430 |  | Independent Apartments - Mallory - 122/430
1b4ccf18-d2af-4052-9981-8ae350d537e1 | section | Details
75e88b9d-b1ac-4340-9c93-28d48d387e94 | checkpoint | Check Shower Pull Cord, as it is failing intermittently (Attach Photos)
87c3abfd-8ab2-4947-ba6e-0b63ae3b7e56 | checklist | Independent Apartments - Mallory - 126 |  | Independent Apartments - Mallory - 126
5722d6b5-a39d-418a-88d5-cf4152c2c9dd | section | Details
759c98ad-482b-44c1-a959-8dfb66c145ad | checkpoint | Tidy wireless pendant cables (Attach Photos)
e2a923f1-1c01-4496-a517-31f566cb6afa | checkpoint | Mount receiver onto the wall (Attach Photos)
acfba419-078a-43e6-8565-a8327fcee309 | checkpoint | Change Ensuite Toilet Call Point insert from Green to Red (Attach Photos)
0ae07611-08f7-4909-af8b-414f7906d8e2 | checkpoint | Change Guest Toilet Call Point insert from Green to Red (Attach Photos)
896e7a3d-08af-4230-9de6-075862637837 | checklist | Independent Apartments - Mallory - 130 |  | Independent Apartments - Mallory - 130
6f1d52a0-cc21-45fc-acb4-a845255bec39 | section | Details
f7030ff2-8658-4ae8-9c53-34b9749c64cd | checkpoint | Check Ensuite Shower Pull Cord, as it is failing intermittently (Attach Photos)
422a320c-e217-482b-b95f-6676a99bd904 | checklist | Independent Apartments - Mallory - 221 |  | Independent Apartments - Mallory - 221
05c198a3-8f37-4135-89e5-eefef0c07568 | section | Details
74cf9d67-7b49-4f50-b217-297560608f2c | checkpoint | Change Lounge Call Point insert from Green to Red (Attach Photos)
a5062153-837a-448b-8f78-4c02e53e6a0d | checkpoint | Change Ensuite Toilet Call Point insert from Green to Red (Attach Photos)
0a0e1da5-8ec0-4857-8141-bb264282acb8 | checkpoint | Change Guest Toilet Call Point insert from Green to Red (Attach Photos)
fa7c3766-bad3-487a-81b1-51a6e0f6f726 | checklist | Independent Apartments - Mallory - 226 |  | Independent Apartments - Mallory - 226
33ad0bb9-d97f-47bf-9c05-e3071d0a9696 | section | Details
00e5ee5f-f06f-47fc-be8a-09e61badf571 | checkpoint | Change Guest Bedroom Call Point insert from Green to Red (Attach Photos)
640179ce-219b-4c7c-9be2-1b3b72bac13f | checkpoint | Change Guest WC Call Point insert from Green to Red (Attach Photos)
935b3d54-7c04-4c0f-8ef8-a9015d9ee438 | checklist | Independent Apartments - Mallory - 230 |  | Independent Apartments - Mallory - 230
5410f6e9-6d32-4a2c-b8b9-857878a4d4d8 | section | Details
b1c8c670-1a60-47ce-aa4e-366551e4e1fb | checkpoint | Check Ensuite Shower Pull Cord reassurance LED, as it's not working when the alarm is active (Attach Photos)
335487d1-99da-47d8-ac1c-6131ee257b9f | checkpoint | Check Master Bedroom Bedhead Pendant removal alarm, as it failed (Attach Photos)
179f8857-7db5-4a5d-b6fc-1825b97a610b | checklist | Independent Apartments - Mallory - 331 |  | Independent Apartments - Mallory - 331
50d7b9f1-4b8b-43fe-853e-75909f1c8b29 | section | Details
740bea8e-4ae8-4d3b-bce5-59e484ad503f | checkpoint | Check Ensuite Shower Pull Cord reassurance LED, as it failed (Attach Photos)
13ae1067-ab76-4f87-a885-c7c2cf3288d3 | checklist | Independent Apartments - Mallory - 334 |  | Independent Apartments - Mallory - 334
c1cc24c5-3d73-4d7a-9686-9f746cb667ff | section | Details
23ee11a9-e001-43ab-acfa-b1fcd7da81de | checkpoint | Check Corded Pendant, as it failed (Attach Photos)
4172b3d4-31cf-4ffa-9a3a-beda7de1b831 | checkpoint | Check Master Bedroom Bedhead Pendant removal alarm, as it failed (Attach Photos)
ca865ff6-8579-47fe-8431-bdd09fbb837b | checklist | Independent Apartments - Mallory - 427 |  | Independent Apartments - Mallory - 427
07bed61f-50af-4fac-8765-fc518ea39892 | section | Details
ff2ffa73-775d-4117-8ef6-3565804c28bf | checkpoint | Change Ensuite Toilet Call Point insert from Green to Red (Attach Photos)
a5e92e07-18e0-4895-afac-3bec4d8f7691 | checkpoint | Change Guest Toilet Call Point insert from Green to Red (Attach Photos)
7112e6fb-dcaa-45da-805a-12ab0f504809 | checkpoint | Change Guest Bedroom Call Point insert from Green to Red (Attach Photos)
cd8c49df-b5d6-449d-b542-3f52d24f7d07 | checklist | Independent Apartments - Mallory - 434 |  | Independent Apartments - Mallory - 434
5747ab45-9577-4a17-8c8e-cb173a43f9e3 | section | Details
9aab4927-a06b-4370-997f-7a1758a607cf | checkpoint | Check Toilet Call Point, as it is not securely mounted (Attach Photos)
b9786ce8-2110-4181-af9f-05644b286bad | checklist | Independent Apartments - Mallory/Tenzing - 321/359 |  | Independent Apartments - Mallory/Tenzing - 321/359
ae229dd0-cfc5-46e4-b599-f086f3611cd0 | section | Details
627b451d-5a08-4da9-b9da-f0865d7245f0 | checkpoint | Change Ensuite Toilet Call Point insert from Green to Red (Attach Photos)
38eb00ab-7fa9-43d6-ac74-35ab9517da63 | checkpoint | Change Guest Toilet Call Point insert from Green to Red (Attach Photos)
7b8028c1-2047-4bb7-be7a-dac52b959cf4 | checklist | Independent Apartments - NC Equipment Cabinet - IA SCOTT - Carpark |  | Independent Apartments - NC Equipment Cabinet - IA SCOTT - Carpark
d1be2a1d-237a-4209-a827-62fc743d2a61 | section | Details
b28ac72e-2a0c-4390-981b-3df5d3ece552 | checkpoint | Replace 12V Battery and label with the replacement date (Attach Photos)
18a5b0ca-94de-417c-9242-e2b15a4bc203 | checkpoint | Clean enclosure and remove junk from floor (Attach Photos)
0c80138c-2d57-427e-8f1a-873176002e49 | checklist | Independent Apartments - Nepal - 301 |  | Independent Apartments - Nepal - 301
6693c3d8-81ed-471b-a159-c668f0a13f60 | section | Details
e86c6c88-4bbd-4392-acd3-cf73940606c5 | checkpoint | Check Shower Pull Cord reassurance LED, as it is not working (Attach Photos)
95072f73-38f8-43fd-ab4a-4a2a0d6752b8 | checklist | Independent Apartments - Nepal - 302 |  | Independent Apartments - Nepal - 302
b19ee907-0441-4f92-b7c6-46ef46fc88e2 | section | Details
ad21acd2-9262-40cb-9003-55598a40263a | checkpoint | Check ODL, as it failed (Attach Photos)
c17baff0-298e-4f9f-9e47-dd53286a4d79 | checkpoint | Change Master Bedroom Bedhead Call Point Faceplate from Green to Red (Attach Photos)
569a8591-6b7e-4c16-8ce5-7953d95ec965 | checkpoint | Change Toilet Call Point Faceplate from Green to Red (Attach Photos)
09684c80-e531-4347-b071-fa81bfc548d4 | checklist | Independent Apartments - Ross - 285 |  | Independent Apartments - Ross - 285
61cefe1b-a3c3-4ad5-8ee6-638ac9c34e89 | section | Details
81a1b059-de2b-40dd-9293-67ac04ec5529 | checkpoint | Check Toilet Call Point, as the cancel button is not working (Attach Photos)
44729ee7-77c4-4bb6-9210-0d0701a26a8c | checklist | Independent Apartments - Scott - 135 |  | Independent Apartments - Scott - 135
f04386a8-8d2b-47d6-8c8b-26ef95560fa1 | section | Details
b388e0e1-77f6-4757-8799-32207eec8a94 | checkpoint | Check Shower Pull Cord, as it is failing intermittently (Attach Photos)
a61a65a7-f50f-49ec-9f85-5f16ae7fefd9 | checkpoint | Change ALL Call Points in the Apartment to have Red inserts (Attach Photos)
65a33836-82bd-45fd-80d9-1379860c0970 | checklist | Independent Apartments - Scott - 136 to 440 |  | Independent Apartments - Scott - 136 to 440
6f2af3aa-1e20-4b03-a367-3e44b8efe659 | section | Details
2df9c8eb-7772-4b80-b880-e26110eb1b08 | checkpoint | Change ALL Call Points in the Apartment to have Red inserts (Attach Photos)
241002ae-1eb2-45f4-b414-3a5a5a469266 | checklist | Independent Apartments - Scott - 235/236 |  | Independent Apartments - Scott - 235/236
783de68b-f995-44eb-a293-7a75cf8b6ac5 | section | Details
871361fa-0962-4c1b-990e-04b6bd61b0f7 | checkpoint | Check Shower Pull Cord, as reassurance LED failed (Attach Photos)
68d257e1-84a7-477c-b68b-c75dd68ed4cd | checkpoint | Change ALL Call Points in the Apartment to have Red inserts (Attach Photos)
cb2a0b8b-a68f-4bbd-a84f-3422b7fe96a7 | checklist | Independent Apartments - Shackleton - 146 |  | Independent Apartments - Shackleton - 146
12905843-5923-4b7c-a6f9-a0f82956385d | section | Details
8b820a53-4610-443d-82fb-19bc6e876897 | checkpoint | Check Guest Bedroom Bedhead Pendant removal alarm, as it failed (Attach Photos)
f73ac1de-ee36-4b2e-ae69-ddbac21596fb | checklist | Independent Apartments - Tenzing - 150/159 |  | Independent Apartments - Tenzing - 150/159
ab67fbdc-060b-4a66-a36a-d00872307f8e | section | Details
dac7cb4a-fe86-4358-8787-17f30efcaf63 | checkpoint | Check Shower Pull Cord reassurance LED, as it failed (Attach Photos)
e61f32db-8ea9-4489-be9f-84f3cb8e15f6 | checklist | Independent Apartments - Tenzing - 160 |  | Independent Apartments - Tenzing - 160
b90a53f1-2529-4de8-b0ff-fd63ff055207 | section | Details
6774c67f-8732-4913-8ed0-e760d415a66e | checkpoint | Check Master Bedroom Bedhead Pendant removal alarm, as it failed (Attach Photos)
30cdbc6e-9646-442a-ae32-14130a5eb42a | checkpoint | Change Toilet Call Point insert from Green to Red (Attach Photos)
ab3916ea-de39-45a6-a6fb-651fe41a75f6 | checklist | Independent Apartments - Tenzing - 347 |  | Independent Apartments - Tenzing - 347
e646b7b0-51ec-4010-918f-91683f93f8cd | section | Details
0c062056-8a54-42b5-8f32-5ae0d3b45779 | checkpoint | Check ODL, as it is failing intermittently (Attach Photos)
2a1c0684-dc1d-470d-a9bb-5a86ef4f2f8f | checklist | Independent Apartments - Tenzing - 453 |  | Independent Apartments - Tenzing - 453
405cca0e-8cd2-407b-be93-5e5432576bf0 | section | Details
a3ab695d-08da-4ab4-80ad-11aa743734f5 | checkpoint | Check Master Bedroom Bedhead Call Point, as the pendent removal alarm failed (Attach Photos)
53bbb1ee-4dc8-4b6b-82f9-0ac85992f1dd | checklist | Independent Apartments - Tenzing/Cook/Ross - 250/452/178/271/372/388 |  | Independent Apartments - Tenzing/Cook/Ross - 250/452/178/271/372/388
6fda83c3-e7ad-4566-bb07-3b090516a1bf | section | Details
57c4e956-13d1-4e5f-952b-a263cc09946f | checkpoint | Check Long Range Wireless Pendant alarm, as it does not display on the Pagers (Attach Photos)
1e65f0c3-8d72-499c-8573-c8cb47bc0015 | checklist | Rest Home Room - 316 |  | Rest Home Room - 316
03f76f9f-6329-4f2c-bea1-04158921dcea | section | Details
cbfe9429-6d82-492c-b1b8-123e4b181110 | checkpoint | Check Nurse Presence Call Point - it's not working and causes the virtual Staff Assist alarm to fail as well (Attach Photos)
4268c566-0b1f-4032-9c38-c1bf68a1af11 | checklist | Rest Home Room - 334 |  | Rest Home Room - 334
69776769-c130-4563-8ee0-4bcf7fe10057 | section | Details
a210b33d-b520-4663-9c2a-7cca72ceea23 | checkpoint | Check Pendant removal alarm, as it fails intermittently (Attach Photos)
5ffcd48b-3562-49d0-854f-79f60920d665 | checklist | Rest Home Room - 335/336 |  | Rest Home Room - 335/336
a765afb7-4fe5-4f7f-85c2-115ce78abae1 | section | Details
b6ea6b8d-406f-4f5c-a3ac-5b760afce85d | checkpoint | Replace Ensuite Toilet Call Point - requires very hard pressing to activate (Attach Photos)
57e2ed0e-7743-4037-b69f-26bc8a6165ff | checklist | Rest Home Room - 338 |  | Rest Home Room - 338
d80a4b4c-01fd-4b19-9186-e83c0a22c8c8 | section | Details
3afab522-d5e6-44e7-948a-3a29ddd3987c | checkpoint | Check Floor Pressure mat, as it failed (Attach Photos)
7a37adf2-857e-40c7-9bdd-0c7bce5c7d36 | checklist | Rest Home Room/Hospital 3 - 314/434/458 |  | Rest Home Room/Hospital 3 - 314/434/458
d4b78377-4238-41ae-aeb0-8dcf35c92e60 | section | Details
d49b4d39-c1dd-4c8e-9ea0-e477f89d16b3 | checkpoint | Install surrounding cover on the Ensuite Toilet Call Point (Attach Photos)
aed2bbc3-78d1-445a-b2ff-06a10f8aa8c7 | checklist | Rest Home Room/Independent Apartments - Tenzing - 329/450 |  | Rest Home Room/Independent Apartments - Tenzing - 329/450
c2bdba85-757f-47fd-92b2-9eac2cd89b5a | section | Details
7f96f819-c81f-456b-932e-abf0b45c8665 | checkpoint | Check Shower Pull Cord, as it failed (Attach Photos)
3e7054ae-a340-4e5c-ac3a-7aeb6bcfd1bb | checklist | Rest Home/Independent Apartments - Shackleton/Cook/Everest/Ross - Lounge 2/341/G77/374/170/191 |  | Rest Home/Independent Apartments - Shackleton/Cook/Everest/Ross - Lounge 2/341/G77/374/170/191
5014c0f3-a853-4fa6-9833-e2273b826492 | section | Details
0675ad66-21df-4638-b30f-143719173a01 | checkpoint | Check ODL, as it failed (Attach Photos)
b8ccd7d6-4f64-4185-bbe4-d2d9736011fb | checklist | SCU - 402/424 |  | SCU - 402/424
fd58564f-bdb1-4070-9bd4-77a724891d91 | section | Details
5249b1e0-d53e-4ae4-82ea-e0cba209918c | checkpoint | Check Pendant removal alarm - activates after a long delay (Attach Photos)
75135b8d-cdea-4aac-bed8-29aa42c48a8f | checklist | SCU - 404 |  | SCU - 404
820158ba-37de-4e67-869c-65acd3a757c0 | section | Details
d7dafb56-b844-40bd-804e-1fc8daf2eef8 | checkpoint | Check Door alarm - activates after a long delay (30 Seconds) when DCS is active (Attach Photos)
e9562301-b360-4346-bd69-7898c69d7a86 | checklist | SCU - 405 |  | SCU - 405
5afa1f06-2013-4032-b969-e6e7f33920c8 | section | Details
2233a46c-dd6f-450e-918b-4cea36e78d11 | checkpoint | Check DCS - failed to track movements inside the room (Attach Photos)
1c2b3369-de12-4fad-af7b-010592f33d1a | checklist | SCU - 406 |  | SCU - 406
45d3bf67-3f6c-4342-9512-3de58ee30e40 | section | Details
1b38076a-4681-40d8-bdd3-e0d2915adfcf | checkpoint | Bedroom light is not working when DCS is active - please check (Attach Photos)
e9da5962-aae5-451e-b7c4-0a9471b09449 | checkpoint | Check DCS panel at door - shows incorrect movements (Attach Photos)
232f6d93-d0ef-4951-a0d0-ff2561bfd915 | checklist | SCU - 408/411 |  | SCU - 408/411
fa09b332-90e0-4b48-92de-b0d78fd3768a | section | Details
0576a787-8d29-4a9b-b28f-6f68771399ca | checkpoint | Install surrounding plate for the bed mat socket (Attach Photos)
181cd8bc-23a6-4d16-a1ee-344825abeda2 | checklist | SCU - 429 |  | SCU - 429
e2ce6efb-d9d8-4d02-951e-5dab7cdf267c | section | Details
77ff7c35-d091-4cab-a7f7-2e2d0c70c74a | checkpoint | Correct Nurse Presence Call Point - it is installed upside down (Attach Photos)
2c9c97e4-d67c-4947-8142-a689e6f99388 | checklist | SCU - Display - SCU Lounge |  | SCU - Display - SCU Lounge
948473c4-c750-499f-a765-a026c6453725 | section | Details
1d0e209a-0517-440f-9f9a-3f7a9df7f170 | checkpoint | Tie up Ethernet connection cable that is hanging below the display (Attach Photos)
86b4f5da-08d7-40f8-bf02-f75d569a5e28 | checkpoint | Clean the display, as it is dirty (Attach Photos)
8230665e-01b6-4021-8e86-db3e6d6adbc6 | checklist | Serviced Apartments - 110 |  | Serviced Apartments - 110
0f5ee500-8bd7-47d7-9621-6fd9c518741f | section | Details
2a513cba-1d69-414d-8bcc-5e6833c0f722 | checkpoint | Check reassurance LED on Bedroom and WC Call Points, very dim and hardly visible (Attach Photos)
590ff985-9a49-4fd7-82cb-cc428370819e | checklist | Serviced Apartments - 112/205 |  | Serviced Apartments - 112/205
56b6f22c-e27c-48ac-8402-61d3f5bf0567 | section | Details
e1aba327-6de5-4fb7-9bcc-5c5c25283954 | checkpoint | Check Staff Assist alarm, it's not working (Attach Photos)
b4bc08e0-eea2-4c21-9453-b2074cb37c3c | checklist | Serviced Apartments - 114/308 |  | Serviced Apartments - 114/308
211e8a8d-f895-447d-af1a-75803170a9a1 | section | Details
2f41e901-eaa9-4e54-afd1-324a4df826f4 | checkpoint | Check Shower Pull Cord reassurance LED, it's not working (Attach Photos)
59880b26-51ce-4b50-b236-7a810ac9484f | checklist | Serviced Apartments - 204 |  | Serviced Apartments - 204
7212bb21-2b27-4cab-ae2a-877f41a7ccb5 | section | Details
525d3efa-856e-4a0c-831a-d485696fa518 | checkpoint | Check Short-Range wireless pendant, as it failed (Attach Photos)
b73a7598-462b-4299-a32f-e6b1b5e30ca2 | checklist | Serviced Apartments - 206/210 |  | Serviced Apartments - 206/210
294094cd-3121-4b4d-b4b4-4230a6e3d5e0 | section | Details
093f8f3e-6a5c-4c81-93e4-cf147c5e8932 | checkpoint | Check Alarm generated by Bedhead Call Point - generating Resident Call when it should be Emergency Alarm (Attach Photos)
2d87cfd4-b310-4e40-a2c2-56b2b7413a6a | checklist | Serviced Apartments - 315 |  | Serviced Apartments - 315
4db067c9-14ef-40fd-92ef-f61a67fe1670 | section | Details
2d23218d-3b3b-4f2b-9d66-6a5dba2a8132 | checkpoint | Check Shower Pull Cord, it gets stuck during pull test and needs to be manually pushed back to reset the alarm (Attach Photos)
74778371-2c44-4e88-befc-52d43a9fcef8 | checklist | Serviced Apartments - 320 |  | Serviced Apartments - 320
5c9fa248-1fc7-45a7-8aa8-2e1936573947 | section | Details
fbb99ea0-bcef-4237-bb27-23105a9f6099 | checkpoint | Check Shower Pull Cord, as it failed (Attach Photos)
d4bdb446-f924-4627-b3b0-07a02203ddd0 | checkpoint | Check Short-Range pendant, as it fails intermittently (Attach Photos)
14adf533-0b6d-44c6-85bd-800171493a18 | checklist | Serviced Apartments/Independent Apartments Everest - 318/266 |  | Serviced Apartments/Independent Apartments Everest - 318/266
f3ad18d0-c31f-40b2-ac69-5f6403fd83e3 | section | Details
c943db08-ba38-43d4-b210-3cbf55622e7e | checkpoint | Check ODL - brightness is very low and hardly visible (Attach Photos)
13cdcf8e-2d09-46fd-9af9-1dbe54aaa2ca | checklist | Townhouse - 2 |  | Townhouse - 2
14bef440-cfdd-4fd9-9ca0-84e3c75e3b73 | section | Details
2342cce1-50ce-447f-955e-d1650147d833 | checkpoint | Change Guest Toilet Call Point Faceplate from Green to Red (Attach Photos)
4eb918b6-1a2f-4ecf-8411-f145668858ed | checklist | Townhouse - 3/4/37/39 |  | Townhouse - 3/4/37/39
cd7f8148-9d58-41e2-88bb-a5f955daa2be | section | Details
85d2ffd4-5377-4ebe-adcd-a8b8edb2ee61 | checkpoint | Change Ensuite Toilet Call Point insert from Green to Red (Attach Photos)
c3ca2b5a-68c4-4ab6-9778-a0e8988c9008 | checklist | Townhouse - 6 |  | Townhouse - 6
f7ff6da8-81d3-4022-b736-a40791d00203 | section | Details
249cd641-ec1d-44e6-8a61-1f87f18fed65 | checkpoint | Change Master Bedroom Call Point insert, should be Red (Attach Photos)
505444e5-6005-48d9-82c1-d75434e72db9 | checklist | Townhouse - 7 |  | Townhouse - 7
83b3e396-02f5-4152-99b4-4644cd46e472 | section | Details
8b6eccad-aef0-47cf-bf9a-2bc1c9a0551e | checkpoint | Change Guest Bedroom Call Point Faceplate to a Red Emergency Call Faceplate (Attach Photos)
a4f2bf3b-6946-4450-a553-1aa07e584b2c | checkpoint | Change Ensuite Toilet Call Point Faceplate to a Red Emergency Call Faceplate (Attach Photos)
041117f9-0cab-459c-85f1-f1d367f4250c | checkpoint | Change Guest Toilet Call Point Faceplate to a Red Emergency Call Faceplate (Attach Photos)
9c5ea8e4-ab1e-4847-9243-3e5e4424762f | checklist | Townhouse - 9 |  | Townhouse - 9
906601b4-d339-40ed-9535-a6a83fb94d84 | section | Details
802ab0b4-01ea-431b-9739-d344c51166cf | checkpoint | Check Master Bedroom Call Point - requires hard pressing to activate (Attach Photos)
86a01bd6-8c33-4329-8b38-7573941506f0 | checkpoint | Check Shower Pull Cord reassurance LED, as it failed (Attach Photos)
1ff2a44e-3444-4cc2-afb9-b2b44d7c3fbe | checkpoint | Change Ensuite Toilet Call Point insert from Green to Red (Attach Photos)
2557dfb7-0e68-4c84-8735-5f108456928b | checkpoint | Change Guest Toilet Call Point insert from Green to Red (Attach Photos)
15ec3615-cf00-4192-82c5-e371141201a0 | checklist | Townhouse - 10 |  | Townhouse - 10
e17a1317-27bb-4c77-8149-536c2ddf4744 | section | Details
d296d22b-f73a-438e-9664-06af3897ced2 | checkpoint | Check Guest Toilet Call Point - requires hard pressing to activate (Attach Photos)
e095bdfd-cd69-49cf-b5f5-2651fa467e6d | checkpoint | Change Ensuite Toilet Call Point insert from Green to Red (Attach Photos)
3750a2ce-02d0-4c33-82b3-550f49d7d331 | checkpoint | Change Guest Toilet Call Point insert from Green to Red (Attach Photos)
c3a945e5-16ff-42bd-a8a5-9fafc573e00d | checklist | Townhouse - 11/15 |  | Townhouse - 11/15
5f50f7ef-6197-4848-932e-7af9294b019f | section | Details
1385914b-29c7-4fb6-a092-8dfdbf571f71 | checkpoint | Change Master Bedroom Bedhead Call Point insert from Green to Red (Attach Photos)
17aa2037-cee1-45b0-a7d4-2ae080e24658 | checkpoint | Change Ensuite Toilet Call Point insert from Green to Red (Attach Photos)
d49147a4-4bb6-4ebf-8068-b9bd290e9907 | checkpoint | Change Guest Toilet Call Point insert from Green to Red (Attach Photos)
4027666b-5254-46ef-94a4-b7711b0ae356 | checklist | Townhouse - 12A/16 |  | Townhouse - 12A/16
db09972b-04da-4238-8132-949691788931 | section | Details
36176c1f-49d9-425a-a3d4-0a153e22cb39 | checkpoint | Change Guest Toilet Call Point insert from Green to Red (Attach Photos)
5a5fe173-f03d-48e0-9a04-2a30d89d7e53 | checklist | Townhouse - 27 |  | Townhouse - 27
3b911758-99f9-4e83-85a4-7a844e156782 | section | Details
3055eda7-c4ef-4df0-8a51-4c300ad328d3 | checkpoint | Replace Call Point in Ensuite Toilet - requires hard press to activate (Attach Photos)
eba5b786-6aaf-4d6d-8f9d-8136cde3fe73 | checklist | Townhouse - 33 |  | Townhouse - 33
7f695452-93b3-4ddd-ad86-1980588bb62f | section | Details
292707ec-08d0-46ce-8752-07003e1dcb94 | checkpoint | Check Smoke Detector (HM) in the Hallway, as it failed the Smoke Test (Attach Photos)
9de03a8f-dbdc-4d33-91df-7fcad1067155 | checkpoint | Check Master Bedroom Bedhead Call Point - requires hard pressing to activate (Attach Photos)
b44636e2-dd77-4a9a-be93-481b430cf599 | checklist | Townhouse - 35 |  | Townhouse - 35
bb59612a-53aa-4b49-a577-c9b5401981ca | section | Details
d6fbb07a-84d0-46d1-a08f-de3762f85acd | checkpoint | Replace battery on Short Range Wireless Pendant - still working, but activation LED on pendant is dim (Attach Photos)
d20a1f3d-d5c7-4f6d-a237-8749a4b67612 | checklist | Townhouse - 36 |  | Townhouse - 36
26060089-7f3c-4039-94f4-25a72453e5d2 | section | Details
95e642e1-6217-4c7b-bf8c-8381f2d096a7 | checkpoint | Check Reassurance Light in Master Bedroom Bedhead Call Point, failed when alarm is active (Attach Photos)
2d18dfa9-81fc-410a-ab17-02571e5c188c | checklist | Townhouse - 43 |  | Townhouse - 43
21df5d7b-bc59-4933-9640-6e97430c3fb7 | section | Details
a141437c-ec0d-437a-912d-43ea37cd24bc | checkpoint | Check Ensuite Shower Pull Cord, it fails intermittently (Attach Photos)
70a79473-4f10-47e3-b837-ee8e30dbeef6 | checklist | Townhouse - 63 |  | Townhouse - 63
6abc3ffa-bdb4-44fe-a56e-1cec3ff35b78 | section | Details
f74e73de-6e49-48dd-9f54-367b563dc5b7 | checkpoint | Check Lounge Smoke Detector (System Sensor), as it failed to activate the alarm on Nurse Call Displays and Pagers (Attach Photos)
4b908e6a-0eb3-41e0-ac48-86c99f0ec601 | checklist | Townhouse - 66 |  | Townhouse - 66
47a0a61a-3698-491b-9270-353c22a36549 | section | Details
90395b14-a291-43c3-887e-86fc4cc25838 | checkpoint | Change Master Bedroom Bedhead Call Point insert from Green to Red (Attach Photos)
c1e6f452-da12-4c5e-8739-86018ea4ae88 | checklist | Townhouse - 79/80 |  | Townhouse - 79/80
a017fdfa-caba-4af5-8fc2-b80d1d245511 | section | Details
3b4cc5a1-0e86-4e21-9072-503229eb0d60 | checkpoint | Check both Smoke Detectors (System Sensor), as they both failed to activate the alarm on Nurse Call Displays and Pagers (Attach Photos)
aa27cc4c-e05d-45a5-b4e1-290a33ee2dee | checklist | Townhouse - 83/84 |  | Townhouse - 83/84
1c0fa0ba-a251-43af-9aca-8b0a05a5629f | section | Details
e59a03b2-6026-4f6d-8b40-3ea3d054611c | checkpoint | Check both Smoke Detectors (VST), as they both failed to activate the alarm on Nurse Call Displays and Pagers (Attach Photos)
ada1dd89-ea93-4627-814d-3010b0fb422a | checklist | Townhouse - 85 |  | Townhouse - 85
0f9001ae-79f1-48f1-a379-e5d273afc27b | section | Details
473bf380-3b90-4687-83e3-eec5c2731ebd | checkpoint | Check both Smoke Detectors (HM), as they both failed to activate the alarm on Nurse Call Displays and Pagers (Attach Photos)
043714a2-d0b8-449f-a7e1-fc0a73382573 | checklist | Townhouse - 87 |  | Townhouse - 87
7f2ce1a4-5d16-4a15-8e28-d84281329db1 | section | Details
0f02c5c8-1bee-4990-a043-a427e1f85f9d | checkpoint | Check both Smoke Detectors (VST in Hallway and HM in the Lounge), as they both failed to activate the alarm on Nurse Call Displays and Pagers (Attach Photos)
2a251586-b377-4e1b-97cf-7fedc1fbf45d | checkpoint | Check Pendant removal alarm, as it failed (Attach Photos)
33e859d5-86ee-49eb-9415-a8a408632f9f | checklist | Townhouse - 90 |  | Townhouse - 90
08da1512-1a7e-4ece-ab19-8983f0591673 | section | Details
5a51cdd5-c4c4-4700-8509-f69a3fcfc07b | checkpoint | Check both Smoke Detectors (System Sensors in the Hallway and Lounge), as they both failed to activate the alarm on Nurse Call Displays and Pagers (Attach Photos)
961043f9-aa74-4fea-b28a-b8251d35434e | checklist | Townhouse/Independent Apartments - Mallory - 56/67/332 |  | Townhouse - 56/67
b14fb2ed-44bf-4032-862c-5ab09451a98d | section | Details
6686c18d-cb24-42ae-be38-f698e39f4cc3 | checkpoint | Check Ensuite Shower Pull Cord, as it failed (Attach Photos)
85dda116-b9cd-4fd1-b77b-972479178d27 | checklist | Townhouse/Independent Apartments - Mallory - 61/71/128 |  | Townhouse/Independent Apartments - Mallory - 61/71/128
feb6d7c4-824f-4f7e-a740-6069aa3b2515 | section | Details
90b07cef-b96c-4863-9b31-5e19ea60cc14 | checkpoint | Check Ensuite Shower Pull Cord reassurance LED, as it's not working when the alarm is active (Attach Photos)
28b3ef7c-d6c2-4605-a6a2-46cd9d94bd0d | checklist | Village Centre - Gym |  | Village Centre - Gym
7dcb0c60-c68c-4c66-92df-b42a6a500b30 | section | Details
ed2e3a14-7e7a-417d-bbbb-a0f56b6ebeb4 | checkpoint | Connect Gym Call Point to the ODL at the Gym Entrance - inside the same area, so should work (Attach Photos)
682b2add-70b0-4127-9505-028b31f29462 | checklist | Village Centre - Gym F WC |  | Village Centre - Gym F WC
dae37ede-073c-4b78-9f39-7faaa014ce8d | section | Details
12d5901b-f248-4369-a979-3aa1ecf7d48c | checkpoint | Check ODL, as it failed (Attach Photos)
25824657-5b3d-4371-a42c-afdf4824db99 | checkpoint | Install ODL above Female Toilet Door (Attach Photos)
3da2df75-ad33-4300-952a-7a65070874bc | checkpoint | Change Call Point insert from green to red (Attach Photos)
4be9acbc-d61c-4fa6-bb3f-baec128032eb | checklist | Village Centre - Lounge, by Bar area |  | Village Centre - Lounge, by Bar area
ac0d7554-71a7-49d1-ae8f-68e181c4fcc6 | section | Details
090aeb82-efd2-4bb3-8ed9-32aa7085e8a7 | checkpoint | Check Call Point by the Bar, as it failed (Attach Photos)
2a8a9451-6bd9-4c13-9ed4-7cda32b1f7c4 | checklist | Village Centre - Pool |  | Village Centre - Pool
45085755-ea8d-4ef8-bf2e-177662c27fad | section | Details
8092ac94-deb5-4549-a9d8-3ff177a04763 | checkpoint | Check Call Point - requires hard pressing to activate (Attach Photos)
3c896587-9cc4-4c4e-9008-24cd51285322 | checklist | Village Centre - SA Main WC Women |  | Village Centre - SA Main WC Women
2affd922-7dab-4ce1-a773-8bd736af5e53 | section | Details
6815e147-f8b1-45aa-8b07-effd5d61971b | checkpoint | Check Call Point - Failed to generate any alarm, should be an Emergency Call Alarm (Attach Photos)
f12ac045-ad1b-4e98-a563-f35264a90989 | checklist | Village Centre - Sales Office |  | Village Centre - Sales Office
57df8904-780b-4325-82dd-df7c82316e39 | section | Details
cc59193f-ded4-40eb-8e01-dc375522e109 | checkpoint | Displays as "Doctors Office", this needs to be updated (Attach Photos)
ca3c1b03-a2fd-4d77-a444-23e21ff8293c | checkpoint | Install surrounding cover on the ODL (Attach Photos)