## res
type | name | values | prefix | postfix | usage count | user addable | template name | Resource Id
checklist | Entire Village - Fire to Nurse Call interface test | [5 empty] | Entire Village - Fire to Nurse Call interface test
section | Details
checkpoint | Programme all Annunciators in the Village with text scrolling (Attach Photos)
checklist | SCU - Display - SCU Lounge | [5 empty] | SCU - Display - SCU Lounge
section | Details
checkpoint | Tie up Ethernet connection cable that is hanging below the display (Attach Photos)
checkpoint | Clean the display, as it is dirty (Attach Photos)
checklist | SCU - 402/424 | [5 empty] | SCU - 402/424
section | Details
checkpoint | Check Pendant removal alarm - activates after a long delay (Attach Photos)
checklist | SCU - 404 | [5 empty] | SCU - 404
section | Details
checkpoint | Check Door alarm - activates after a long delay (30 Seconds) when DCS is active (Attach Photos)
checklist | SCU - 405 | [5 empty] | SCU - 405
section | Details
checkpoint | Check DCS - failed to track movements inside the room (Attach Photos)
checklist | SCU - 406 | [5 empty] | SCU - 406
section | Details
checkpoint | Bedroom light is not working when DCS is active - please check (Attach Photos)
checkpoint | Check DCS panel at door - shows incorrect movements (Attach Photos)
checklist | SCU - 408/411 | [5 empty] | SCU - 408/411
section | Details
checkpoint | Install surrounding plate for the bed mat socket (Attach Photos)
checklist | SCU - 429 | [5 empty] | SCU - 429
section | Details
checkpoint | Correct Nurse Presence Call Point - it is installed upside down (Attach Photos)