## res
type | name | values | prefix | postfix | usage count | user addable | template name | Resource Id
checklist | Hospital - Display - Hospital 2 - Ollivier
section | Details
checkpoint | Check Displays in the Hallway near Room 228, one looks like it has no power (Attach Photos)
checklist | Hospital - Hospital 2 - Room 203
section | Details
checkpoint | Check Staff Assist alarm, text description is wrong on the displays - shows "H208" (Attach Photos)
checkpoint | Check Floor Pressure mat, as it failed (Attach Photos)
checklist | Hospital - Hospital 2/3 - Room 207/208/220/451
section | Details
checkpoint | Install surrounding cover on the Bedhead Call Point (Attach Photos)
checklist | Hospital - Hospital 2 - Room 210
section | Details
checkpoint | Check Staff Assist Alarm, as it failed (Attach Photos)
checkpoint | Check Floor Pressure mat, as it failed (Attach Photos)
checklist | Hospital - Hospital 2 - Room 212/219
section | Details
checkpoint | Replace Bedhead Call Point - requires very hard pressing to activate (Attach Photos)
checklist | Hospital - Hospital 2/3/Independent Apartments - Tenzing - Room 221/238/435/442/447/452/455
section | Details
checkpoint | Tidy wireless pendant cables (Attach Photos)
checkpoint | Mount receiver onto the wall (Attach Photos)
checklist | Hospital - Hospital 2 - Room 223/242
section | Details
checkpoint | Check wired pendant lead, as it failed (Attach Photos)
checklist | Hospital - Hospital 3 - Room 434/458
section | Details
checkpoint | Install surrounding cover on the Ensuite Toilet Call Point (Attach Photos)
checklist | Hospital - Hospital 3 - Room 454
section | Details
checkpoint | Check Staff Assist alarm, text description is wrong on the displays - shows "H3 Rm455" (Attach Photos)
checkpoint | Tidy wireless pendant cables (Attach Photos)
checkpoint | Mount receiver onto the wall (Attach Photos)
checklist | Hospital - Hospital 3 - Room 457
section | Details
checkpoint | Replace insert in the Ensuite Toilet Call Point, requires Green Insert (Attach Photos)
checklist | Hospital - Hospital 3 - Room 459
section | Details
checkpoint | Install surrounding cover on the Bedhead Call Point (Attach Photos)
checkpoint | Tidy wireless pendant cables (Attach Photos)
checkpoint | Mount receiver onto the wall (Attach Photos)
checklist | Hospital - Hospital 3 - Toilet
section | Details
checkpoint | Install surrounding cover on the Toilet Call Point (Attach Photos)