## res
type | name | values | prefix | postfix | usage count | user addable | template name | Resource Id
checklist | Independent Apartment - Mallory/Shackleton - 121/433/146
section | Details
checkpoint | Replace Shower Pull Cord, as it requires to be manually pushed back up to rest position to cancel the call (Attach Photos)
checklist | Independent Apartments - Mallory - 122/430
section | Details
checkpoint | Check Shower Pull Cord, as it is failing intermittently (Attach Photos)
checklist | Independent Apartments - Mallory - 126
section | Details
checkpoint | Tidy wireless pendant cables (Attach Photos)
checkpoint | Mount receiver onto the wall (Attach Photos)
checkpoint | Change Ensuite Toilet Call Point insert from Green to Red (Attach Photos)
checkpoint | Change Guest Toilet Call Point insert from Green to Red (Attach Photos)
checklist | Independent Apartments - Mallory - 130
section | Details
checkpoint | Check Ensuite Shower Pull Cord, as it is failing intermittently (Attach Photos)
checklist | Independent Apartments - Mallory - 221
section | Details
checkpoint | Change Lounge Call Point insert from Green to Red (Attach Photos)
checkpoint | Change Ensuite Toilet Call Point insert from Green to Red (Attach Photos)
checkpoint | Change Guest Toilet Call Point insert from Green to Red (Attach Photos)
checklist | Independent Apartments - Mallory - 226
section | Details
checkpoint | Change Guest Bedroom Call Point insert from Green to Red (Attach Photos)
checkpoint | Change Guest WC Call Point insert from Green to Red (Attach Photos)
checklist | Independent Apartments - Mallory - 230
section | Details
checkpoint | Check Ensuite Shower Pull Cord reassurance LED, as it's not working when the alarm is active (Attach Photos)
checkpoint | Check Master Bedroom Bedhead Pendant removal alarm, as it failed (Attach Photos)
checklist | Independent Apartments - Mallory/Tenzing - 321/359
section | Details
checkpoint | Change Ensuite Toilet Call Point insert from Green to Red (Attach Photos)
checkpoint | Change Guest Toilet Call Point insert from Green to Red (Attach Photos)
checklist | Independent Apartments - Mallory - 331
section | Details
checkpoint | Check Ensuite Shower Pull Cord reassurance LED, as it failed (Attach Photos)
checklist | Independent Apartments - Mallory - 334
section | Details
checkpoint | Check Corded Pendant, as it failed (Attach Photos)
checkpoint | Check Master Bedroom Bedhead Pendant removal alarm, as it failed (Attach Photos)
checklist | Independent Apartments - Mallory - 427
section | Details
checkpoint | Change Ensuite Toilet Call Point insert from Green to Red (Attach Photos)
checkpoint | Change Guest Toilet Call Point insert from Green to Red (Attach Photos)
checkpoint | Change Guest Bedroom Call Point insert from Green to Red (Attach Photos)
checklist | Independent Apartments - Mallory - 434
section | Details
checkpoint | Check Toilet Call Point, as it is not securely mounted (Attach Photos)
checklist | Independent Apartments - NC Equipment Cabinet - IA SCOTT - Carpark
section | Details
checkpoint | Replace 12V Battery and label with the replacement date (Attach Photos)
checkpoint | Clean enclosure and remove junk from floor (Attach Photos)
checklist | Independent Apartments - Scott - 135
section | Details
checkpoint | Check Shower Pull Cord, as it is failing intermittently (Attach Photos)
checkpoint | Change ALL Call Points in the Apartment to have Red inserts (Attach Photos)
checklist | Independent Apartments - Scott - 235/236
section | Details
checkpoint | Check Shower Pull Cord, as reassurance LED failed (Attach Photos)
checkpoint | Change ALL Call Points in the Apartment to have Red inserts (Attach Photos)
checklist | Independent Apartments - Scott - 136 to 440
section | Details
checkpoint | Change ALL Call Points in the Apartment to have Red inserts (Attach Photos)
checklist | Independent Apartments - Shackleton - 146
section | Details
checkpoint | Check Guest Bedroom Bedhead Pendant removal alarm, as it failed (Attach Photos)
checklist | Independent Apartments - Tenzing - 150/159
section | Details
checkpoint | Check Shower Pull Cord reassurance LED, as it failed (Attach Photos)
checklist | Independent Apartments - Tenzing - 160
section | Details
checkpoint | Check Master Bedroom Bedhead Pendant removal alarm, as it failed (Attach Photos)
checkpoint | Change Toilet Call Point insert from Green to Red (Attach Photos)
checklist | Independent Apartments - Tenzing/Cook/Ross - 250/452/178/271/372/388
section | Details
checkpoint | Check Long Range Wireless Pendant alarm, as it does not display on the Pagers (Attach Photos)
checklist | Independent Apartments - Tenzing - 347
section | Details
checkpoint | Check ODL, as it is failing intermittently (Attach Photos)
checklist | Independent Apartments - Tenzing - 453
section | Details
checkpoint | Check Master Bedroom Bedhead Call Point, as the pendent removal alarm failed (Attach Photos)
checklist | Independent Apartments - Cook - Activities Female WC
section | Details
checkpoint | Check both entrances ODL's, as one failed (Attach Photos)
checklist | Independent Apartments - Cook - 175
section | Details
checkpoint | Remove the Miracle RF Wireless Pendant Receiver that is installed, as the resident does not have a wireless pendant (Attach Photos)
checklist | Independent Apartments - Cook - 176
section | Details
checkpoint | Replace Call Socket on the Master Bedroom Bedhead Call Point, as It is broken (Attach Photos)
checklist | Independent Apartments - Cook - 279
section | Details
checkpoint | Provide longer corded pendant, as resident advises current 1.5m pendant only just reaches across (Attach Photos)
checklist | Independent Apartments - Everest - 163
section | Details
checkpoint | Check Master Bedroom Bedhead Pendant removal alarm, as it failed (Attach Photos)
checklist | Independent Apartments - Everest - 263
section | Details
checkpoint | Check Master Bedroom Bedhead Pendant removal alarm, as it failed (Attach Photos)
checkpoint | Check Guest Showers Pull Cord Call Point, as it is not securely mounted (Attach Photos)
checklist | Independent Apartments - Ross - 285
section | Details
checkpoint | Check Toilet Call Point, as the cancel button is not working (Attach Photos)
checklist | Independent Apartments - Nepal - 301
section | Details
checkpoint | Check Shower Pull Cord reassurance LED, as it is not working (Attach Photos)
checklist | Independent Apartments - Nepal - 302
section | Details
checkpoint | Check ODL, as it failed (Attach Photos)
checkpoint | Change Master Bedroom Bedhead Call Point Faceplate from Green to Red (Attach Photos)
checkpoint | Change Toilet Call Point Faceplate from Green to Red (Attach Photos)